## Sheet1
ENDOSCOPE REPROCESSING AUDIT
Facility Name:
Facility #:
Review Date:
Reviewer:
ENDOSCOPES
Current # of scopes:
Brand: (dropdown)
Generations:
(190’s, etc):
Model #’s
Recent repairs/trends:
Maintenance contract? (dropdown)
Repair log present? (dropdown)
Attach scope inventory list (dropdown)
MANUAL CLEANING
ENZYMATIC CLEANER:
DILUTION RATIO:
RECOMMENDED TEMPERATURE: (dropdown)
EXPOSURE REQUIREMENT: (dropdown)
CHANNEL BRUSHES: (dropdown)
VALVE BRUSHES: (dropdown)
AUTOMATED ENDOSCOPE REPROCESSOR (AER)
BRAND:
MAKE AND MODEL:
NUMBER OF THESE AER MACHINES:
HIGH‐LEVEL DISINFECTANT:
RINSE REQUIREMENTS:
RECOMMENDED TEMPERATURE:
EXPOSURE TIME REQUIREMENT:
MAXIMUM REUSE PERIOD:
Open Container shelf life:
MEC TEST STRIPS:
DIP TIME:
READ TIME:
Is the HLD approved for use in the AER? (dropdown)
IF ANOTHER BRAND/MODEL IS USED:
BRAND:
MAKE AND MODEL:
NUMBER OF THESE AER MACHINES:
HIGH‐LEVEL DISINFECTANT:
RINSE REQUIREMENTS
RECOMMENDED TEMPERATURE:
EXPOSURE TIME REQUIREMENT:
MAXIMUM REUSE PERIOD:
Open Container shelf life:
MEC TEST STRIPS:
DIP TIME:
READ TIME:
Are the AER hook-ups the correct part for each scope accessory? (dropdown)
|  | Audit Step | [5 empty] | Yes | No | Notes/Issues/Comments
|  | I. |  |  |  | SCOPE CLEANING PROCESS
|  |  | 1) |  |  | Review the pre‐cleaning process:
|  |  |  | a |  | Angulation knobs unlocked and variable stiffness scopes are turned to the most flexible position before beginning bedside pre‐cleaning | a |  |  | 1Note: Refer to brand specific instructions for appropriate ratio of enzymatic to water and temperature
|  |  |  | b |  | Enzymatic detergents (if used) are diluted1 and freshly prepared2 for one‐time use /or enzymatic sponge used as directed with each scope and discarded after each patient use. ____(Not applicable) | b |  |  | 2Note: Intercept ™can be mixed for 1 day use
|  |  |  | c |  | At the bedside when procedure is complete, soiled endoscopes are wiped and channels are flushed with water (Olympus only) or enzymatic detergent solution immediately after use | c
|  |  |  | d |  | Air/water channel flushed as directed (Olympus/Pentax/Fuji) | d
|  |  |  | e |  | Auxiliary channel is flushed with flushing pump/syringe ____(Not applicable) | e
|  |  |  | f |  | Scopes are transported in covered, impervious containers (not chux) to the re‐processing area, – labeled as Biohazardous and cleaned with a disinfectant wipe after each use | f
|  |  | 2) |  |  | Review the manual scope cleaning process: |  | Yes | No | Notes/Issues/Comments
|  |  |  | a |  | Impervious gown, gloves and face shield are worn during the entire cleaning process | a
|  |  |  | b |  | After valves are removed, the leak tester is attached and the scope is pressurized before submerging it completely in plain water | b |  |  | Full‐strength enzymatic solution is never used directly on sponge or scopes
|  |  |  | c |  | The entire scope is closely observed for leaks while flexing the distal portion in all directions and depressing the video buttons | c
|  |  |  | d |  | The leak tester is then removed (The leak tester is never attached or removed underwater) | d
|  |  |  | e |  | There is a procedure to follow to maintain positive scope pressure during the cleaning and disinfection processes when a leak is detected to prevent water damage | e
|  |  |  | f |  | There is a back‐up leak tester as needed | f
|  |  |  | g |  | Enzymatic detergent (medical grade, low‐foaming, neutral pH detergent formulated for flexible endoscopes) is added to the water‐ verify dilution and water temperature1 | g
|  |  |  | h |  | Center has an enzymatic detergent (or Dawn if approved by scope manufacturer) to use for scopes coated in lipids (i.e. Ruhof Endozyme SLR) | h
|  |  |  | i |  | Scopes are washed using a lint‐free cloth or sponge (enzymatic sponges are used as directed) | i
|  |  |  | j |  | All channels are brushed until no debris visible on the brush, and then flush with enzymatic solution. | j
|  |  |  | k |  | If the scope has an auxiliary water channel (EGD or Colonoscope), it is flushed whether or not a flushing pump/ or the auxiliary water channel was used | k
|  |  |  | l |  | Single‐use brushes and sponges are opened just prior to use and discarded after cleaning of one scope‐ they are never re‐used. Reusable brushes are HLD or sterilized after each use as directed. | l
|  |  |  | m |  | Scopes are exposed to enzymatic cleaner for appropriate time (if specified) | m
|  |  |  | n |  | The scope and all channels are rinsed and flushed with plain water, then purged with air | n
|  |  |  | o |  | Valves are manually washed, brushed and rinsed with water | o
|  |  |  | p |  | A visual inspection for cleanliness is performed after rinsing | p
|  |  |  | q |  | All scopes are cleaned in the same manner regardless of patient’s medical history (i.e. HIV, Hep. B/C) | q
|  | II. |  |  |  | HIGH‐LEVEL DISINFECTION PROCESS
|  |  | 1) |  |  | For Automated Endoscope Reprocessing, confirm compliance with the following: |  | Yes | No
|  |  |  | a |  | Proper PPE is worn (¨gown ¨gloves ¨facemask with shield ¨goggles ) | a
|  |  |  | b |  | HLD is properly prepared (if mixing is required) | b
|  |  |  | c |  | Each disinfectant reservoir is clearly labeled with an ‘in‐use’ and ‘expiration’ date, accurately calculated | c
|  |  |  | d |  | HLD Temperature is set correctly for high‐level disinfectant used | d
|  |  |  | e |  | Disinfectant soak time cycle is set correctly for high‐level disinfectant used | e
|  |  |  | f |  | HLD MEC is tested prior to each and every use with HLD brand‐specific test strips by the staff member that is loading the scope in the AER and initiating the HLD process. | f
[5 empty] | 1 | Correct test strip dip‐time and read‐time is followed using a timer for accuracy | g
[5 empty] | 2 | MEC testing results are documented at time of testing and prior to initiation of endoscope high level disinfection." | h
|  |  |  | g |  | Alcohol flush is performed by AER and bottle is refilled as needed/ or performed manually | g
|  |  |  | h |  | Cycle completion is verified and documented on printout or log before scope is removed from AER | h
[5 empty] | 1 | If cycle completion cannot be verified, the entire disinfection cycle including MEC is repeated | 1
|  |  |  | i |  | Scopes are used immediately or hung in scope cabinet tor complete drying | i
|  |  |  | j |  | AER Machine printouts /manual documentation contain the following: | j
[5 empty] | 1 | Date & time of cycle | 1
[5 empty] | 2 | Scope Identifier (serial number) | 2
[5 empty] | 3 | Patient name/Identifier (system is in place when scope was not used on a patient, ‘9999’) | 3
[5 empty] | 4 | Center personnel Identification | 4
[5 empty] | 5 | MD Identifier | 5
[5 empty] | 6 | MEC – Pass/Fail | 6
[5 empty] | 7 | Temperature of HLD | 7
[5 empty] | 8 | Disinfectant soak time | 8
[5 empty] | 9 | Rinse cycle – Confirm use of appropriate number of rinses after HLD | 9
[5 empty] | 10 | Rinsed endoscopes and channels are then purged with filtered forced air | 10
[5 empty] | 11 | All endoscope channels are flushed with 70% alcohol followed by air purge | 11
[5 empty] | 12 | Initials or signature of personnel verifying cycle: scope #, temp and soak time when removing scope (if documenting manually, the time the scope was removed is documented) | 12
|  |  |  | k |  | Only scopes and other items verified by the AER manufacturer are reprocessed in the AER | k
|  |  |  | l |  | of High‐Level Disinfectant Changes | l
[5 empty] | 1 | A separate log is used for each reservoir | 1
[5 empty] | 2 | Log contains: in‐use date, day of week and time; exp date and day of week and time; reason changed; initials | 2
[5 empty] | 3 | Each log verifies that the HLD is discarded at either the maximum days of reuse and/or when MEC fails (whichever comes first) | 3
|  |  |  | m |  | Filter Changes | m
[5 empty] | 1 | All filter changes are completed at appropriate time frames and documented | 1
[5 empty] | 2 | Water lines are sanitized after internal water filter change and documented in log | 2
|  |  |  | n |  | Preventative Maintenance and Repairs | n
[5 empty] | 1 | A log of repairs and PM’s performed verifies that AER is returned to appropriate settings | 1
[5 empty] | 2 | AER operating manual and the technical support number is immediately available | 2
|  |  | 2) |  |  | For Manual High‐Level Disinfection, confirm compliance with the following: |  | Yes | No
|  |  |  | a |  | Uses proper PPE (gown, gloves and face shield) | a
|  |  |  | b |  | HLD is properly prepared if required | b
|  |  |  | c |  | Each soak pan is clearly labeled with an ‘in‐use’ and ‘expiration’ date, accurately calculated | c
|  |  |  | d |  | HLD temperature is measured prior to each use and within recommended range | d
|  |  |  | e |  | HLD is heated if needed to reach HLD minimum required temperature | e
|  |  |  | f |  | HLD MEC is tested prior to each and every use with HLD brand‐specific test strips | f
[5 empty] | 1 | Correct test strip dip‐time and read‐time is followed using a timer for accuracy | 1
[5 empty] | 2 | MEC testing results are documented on manual soak log | 2
|  |  |  | g |  | Items and reusable medical devices are disassembled, pre-cleaned and completely purged and submersed, with all air bubbles removed, to allow the HLD to make contact with all surfaces. | g
[5 empty] | 1 | If the scope has an auxiliary water channel (EGD or Colonoscope), it is flushed with HLD whether or not a flushing pump/ or the auxiliary water channel was used | 1
[5 empty] | 2 | Items labeled for single‐use are not reprocessed. Discard after use. | 2
[5 empty] | 3 | Items are soaked for the minimum required time for HLD using a timer/clock for accuracy after all channels are purged with high level disinfectant | 3
[5 empty] | 4 | Timing is restarted if additional items are added | 4
|  |  |  | h |  | All items are rinsed/flushed adequately following HLD as per IFU’s (note required # of rinses) | h
Scopes that are manually reprocessed must be rinsed as per IFU, followed by an air purge, then a 70% alcohol flush followed by an air purge, and then dried (alcohol rinse must be separate from water rinse). Any item with a channel must be rinsed per IFU's, followed by an air purge and alcohol flush followed by another air purge, and then dried (alcohol rinse must be separate from water rinse)
|  |  |  | i |  | Scopes are used immediately or hung in scope cabinet for complete drying | i
|  |  |  | j |  | Manual Reprocessing Log Documentation Review: | j
[5 empty] | 1 | Each manual soak is documented when all items are put in the soak pan on the HLD soak log | 1
[5 empty] | 2 | Date; time‐in ; MEC; temp‐in; items soaked; patient name/identifier (for scope reprocessing) and personnel initials recorded | 2
[5 empty] | 3 | Time‐out; temp‐out; and initials of the personnel removing the items is recorded | 3
|  |  |  | k |  | Documentation of High‐Level Disinfectant Changes: | k
[5 empty] | 1 | A separate log is used for each soak pan | 1
[5 empty] | 2 | Log contains: in‐use date, day of week and time; exp date and day of week; reason changed; initials | 2
[5 empty] | 3 | Each log verifies that the HLD is discarded at either the maximum days of reuse and/or when MEC fails (whichever comes first) | 3
|  |  | 3) |  |  | MEC Quality Control is completed and documented: |  | Yes | No
|  |  |  | a |  | QC testing is performed for when each bottle of MEC test strips is opened | a
|  |  |  | b |  | Proper QC procedure is followed (# of strips used, preparation of negative control solution) | b
|  |  |  | c |  | Documentation includes date/time/ brand of test strips/ and initials | c
|  |  |  | d |  | A log of all MEC QC documentation is maintained and verified regularly | d
|  |  |  | e |  | MEC test strip containers are kept closed, labeled with the date opened and the expiration date | e
|  |  | 4) |  |  | 4) Storage After High‐Level Disinfection |  | Yes | No
|  |  |  | a |  | Gloves are worn at all times when handling clean scopes | a
|  |  |  | b |  | Disinfected endoscopes are hung vertically with all valves and caps removed, angulation knobs and stiffening controls in the neutral position | b
|  |  |  | c |  | Scopes are stored in a contained, vented scope cabinet in a clean area – away from dirty reprocessing areas and ends of the scopes are not touching the floor of the scope cabinet. | c
|  |  |  | d |  | The center has a system for determining when scopes were reprocessed so the interval of storage can be determined. | d
|  | III. |  |  |  | Flushing Devices: used during manual scope cleaning, confirm compliance with IFU on the following: |  | Yes | No
|  |  | 1) |  |  | Flow validation test is completed per IFU. Document daily; Flushing device is not to be used if it does not pass Flow Validation testing (Daily Quality Assurance Log) | 1)
|  |  | 2) |  |  | Flushing time is programmed for the brand and model of scope used – see the Rinse Matrix for flushing device; flushing time is set per IFU | 2)
|  |  | 3) |  |  | Scopes are flushed with enzymatic solution, water and air | 3) |  |  | Device Used: (dropdown)
|  |  | 4) |  |  | Decontamination procedure is completed and documented based on the method of fluid draw or IFU. | 4)
|  |  |  | a |  | Daily (if sink used as enzymatic/water source) | a
|  |  |  | b |  | Monthly ( if a clean container is used for enzymatic/water source) | b
|  |  | 5) |  |  | The Scope Buddy remains primed overnight and weekends with water or disinfectant | 5) |  |  | Programmed Flushing time:
|  |  | 6) |  |  | Check IFU’s for prohibited chemical use (such as: alcohol, bleach and hydrogen peroxide) | 6)
|  |  | 7) |  |  | Only room‐temperature solutions are used with Scope Buddy | 7)
|  | IV. |  |  |  | Flushing Pump used for irrigation during a procedure, confirm compliance with the following: |  | Yes | No
|  |  | 1) |  |  | Single use items used with the flushing pump are discarded after use | 1) |  |  | OFP List Model No:
|  |  | 2) |  |  | Reusable items (i.e. channel adapters, connecting tubing, pump cartridges and irrigation bottles) are manually cleaned and high‐level disinfected or sterilized after each patient use according to IFU. Note: AER units are not approved for processing of most flushing pump accessories. | 2) |  |  | Choose one of the following brands: (dropdown)
|  |  | 3) |  |  | Olympus auxiliary tubing: MAJ 855 - Verify green one‐way (2 wings) valve is in place where filter is attached | 3)
|  |  | 4) |  |  | Only sterile water3 is used with the flushing pump | 4)
|  |  | 5) |  |  | Reusable flushing pump bottle sterilized daily OR 24‐hour disposable tubing, connector and sterile water bottle is labeled and changed per IFU | 5) |  |  | 3Note: Mylicon should not be added to flushing pump bottles/sterile H20 bottles
|  | V. |  |  |  | Auxillary water channels and any reusable Auxiliary Water tube used, (ex. Olympus MAJ 855, Pentax OF B113, Fujinon J cleaning adapter) please confirm the following: (there can be auxiliary water channels on EGD scopes- confirm that all channels are flushed and the proper attachments are hooked up with the AER) |  | Yes | No
|  |  | 1) |  |  | The reusable water tube and/or the auxiliary water channel is manually cleaned after every use according to IFU, regardless of whether or not the flushing pump was used during the procedure. | 1) |  |  | 4Note:Not validated AER’s:Custom Ultrasonics System 83 Plus , MV1, CER 1, CER 2, CER Optima, Advantage Plus to reprocess the MAJ‐855 tubing
|  |  | 2) |  |  | After manual cleaning of reusable tubing, follow IFU for either high‐level disinfected or steam‐sterilized | 2)
|  |  | 3) |  |  | If HLD in the AER, the Center has verified that the AER manufacturer has approved the processing of any adapter or scope cleaning device in the AER unit4 | 3)
|  | VI. |  |  |  | Water Bottles3, confirm compliance with the following:
|  |  | 1) |  |  | Reusable water bottles are filled with sterile water3 and changed daily | 1) |  |  | 5Note: Pentax water bottles must be sterilized only
|  |  | 2) |  |  | Reusable water bottles are cleaned daily according to IFU | 2)
|  |  | 3) |  |  | Reusable water bottle lid/tubing is cleaned according to IFU | 3)
|  |  | 4) |  |  | Reusable water bottles are sterilized5 or high‐level disinfected daily according to IFU | 4) |  |  | 6Note: Connector is either labeled single use or 24
|  |  | 5) |  |  | 24‐hour disposable tubing assembly, connector (6) and sterile water bottle is labeled and changed per IFU. | 5)
|  | VII. |  |  |  | Reusable Forceps/ Graspers, and other critical devices (that break the mucosal barrier) must be mechanically cleaned and sterilized between each patient use. Confirm compliance with the following with appropriate sterilizer IFU (7 ): |  | Yes | No
|  |  | 1) |  |  | The center has manufacturer instructions for reprocessing reusable forceps and accessories | 1) |  |  | 7 Note: Refer to Sterilizer audit and competency
|  |  | 2) |  |  | If there is a limited number of times a device can be reprocessed, there is a system to track usage | 2)
|  |  | 3) |  |  | Ultrasonic cleaning is performed on coiled reusable accessories such as biopsy forceps | 3)
|  |  | 4) |  |  | Check IFU for use with milk bath if using movable parts | 4)
|  | VIII. |  |  |  | Reprocessing Room: |  | Yes | No
|  |  | 1) |  |  | All reprocessing is done in an area separate from the procedure room | 1)
|  |  | 2) |  |  | There are clearly identified Clean and Dirty areas in the Reprocessing room | 2)
|  |  | 3) |  |  | There is a one- way flow from dirty to clean to prevent cross-contamination | 3)
|  |  | 4) |  |  | Impervious gowns, gloves, and eye/face protection are available to all performing scope reprocessing | 4)
|  |  | 5) |  |  | There is a sink for hand hygiene or alcohol based hand rub in the reprocessing room | 5)
|  |  | 6) |  |  | Utility sink(s) appropriate to the size and type of endoscope used. ( large enough to allow for total immersion of the scope without being tightly coiled) | 6)
|  |  | 7) |  |  | An operational eye-wash station is available within 20 seconds from the scope reprocessing room | 7)
|  |  | 8) |  |  | SDS information for all scope room chemicals are organized and readily accessible by all personnel | 8)
|  |  | 9) |  |  | Door(s) to the reprocessing room are kept closed at all times | 9)
|  |  | 10) |  |  | The reprocessing room has the required air exchanges and ventilation requirements (check state regs). Reprocessing Room: | 10)
|  |  | 11) |  |  | There is documented, regular maintenance of the air –exchange equipment (e.g. ventilation system, exhaust hoods). Note: After initial air exchange verification or renovations, air exchange is tested periodically or investigation of complaints. Staff completes Monthly room airflow testing. | 11)
|  |  | 12) |  |  | Chemical products are maintained in closed, covered, labelled containers at all times. | 12)
|  |  | 13) |  |  | Lighting is adequate for personnel to observe scope reprocessing steps | 13)
|  |  | 14) |  |  | Center has AER printouts/logs stored in an organized manner (sorted by month/year) | 14)
|  |  | 15) |  |  | Center maintains 3 years of documentation | 15)
|  |  | 16) |  |  | 16) Manufacturer’s instructions are maintained in an organized fashion for the following: | 16)
|  |  |  | a |  | High-Level Disinfectants. Note: Only chemicals approved by AER manufacturer to use. (highlight HLD time, temp and re-use) Note: Custom Ultrasonics use only OPA with re-use period not to exceed 14 day at a minimum of 20 degrees Celsius for 12 minutes. | a
|  |  |  | b |  | Enzymatic Cleaners | b
|  |  |  | c |  | MEC Test Strips (specific to HLD product ) | c
|  |  |  | d |  | Automatic Endoscope Reprocessors | d
|  |  |  | e |  | Scope Flushing Devices | e
|  |  |  | f |  | Leak Testers | f
|  |  |  | g |  | Scopes | g
|  |  |  | h |  | Water Bottles | h
|  |  |  | i |  | Flushing Pumps | i
|  |  |  | j |  | Dilators/Guidewires, and Scope Cleaning Accessories – Reusable | j
|  |  |  | k |  | There is a process to routinely review manufacturers alerts, advisories and technical bulletins relative to the all of the equipment and supplies utilized for reprocessing of endoscopes. | k
|  | IV. |  |  |  | High‐Level Disinfectant Safety |  | Yes | No
|  |  | 1) |  |  | Center has appropriate spill kit for HLD used | 1)
|  |  | 2) |  |  | Center is performing Glutaraldehyde vapor monitoring testing3 (if applicable) | 2)
|  |  | 3) |  |  | Spill drills are performed | 3)
Note: Glutaraldehyde monitoring should be conducted after initiating use of Glutaraldehyde solutions; whenever there is a significant change
in protocol, work practices, caseload, workplace ventilations systems, and after major equipment repairs OR if employees have complaints or symptoms of Glutaraldehyde exposure. Otherwise, routine testing is adequate – annually at a minimum.
|  | X. |  |  |  | Staff Education and Competencies: upon hire and at least annually for the following processes: |  | Yes | No
|  |  | 1) |  |  | Scope Cleaning – Pre‐clean, Manual Cleaning and Storage | 1)
|  |  | 2) |  |  | Manual High ‐Level Disinfection | 2)
|  |  | 3) |  |  | MEC Test strip Quality Controls | 3)
|  |  | 4) |  |  | Testing of HLD with MEC test strip | 4)
|  |  | 5) |  |  | Automated Endoscope Reprocessor | 5)
|  |  | 6) |  |  | Scope Flushing Device: Flow validation, decontamination procedures, flushing times | 6)
|  |  | 7) |  |  | Irrigation Flushing Pump: Setup and use | 7)
|  |  | 8) |  |  | Olympus MAJ‐855 Auxiliary Water tube (if applicable) | 8)
|  |  | 9) |  |  | Reusable Water Bottles: Cleaning and HLD/Sterilization | 9)
|  |  | 10) |  |  | Spill containment plan for all chemicals used in the scope reprocessing room | 10)
Audit deficiencies are identified below. Document the corrective action(s) determined for each deficiency, as well as the responsible party and target date for completion. Post‐ implementation monitoring and follow‐up should be included in the plan of correction.
[6 empty] | ISSUE IDENTIFIED |  | ACTION | RESPONSIBILITY | TARGET DATE
1
2
3
4
5
6
7
8
9
10
11
12
|  | 1 | Attestation: The above reprocessing checklist has been completed by staff interviews, observations and document reviews. The findings and action plan have been reviewed and discussed with the Center Leader/Administrator and a plan for follow‐up has been identified.
SIGNATURE OF CLINICAL DIRECTOR:
DATE:
|  | 2 | Attestation: I have reviewed this reprocessing audit and action plan and will ensure that corrective actions are implemented within the designated time‐ frames. I will notify the Clinical Director immediately if any items on the action plan cannot be completed or if I determine that deviations from the recommended processes have occurred. In addition, I will provide the Clinical Director with a status update on all incomplete items.
SIGNATURE OF CENTER DIRECTOR/ADMINISTRATOR:
DATE:

## LIST to hide (hidden)
Current # of scopes: brand |  | CHANNEL BRUSHES:BRAND | VALVE BRUSHES:BRAND | Maintenance contract?
Olympus |  | DISPOSABLE | DISPOSABLE | Repair log present?
Pentax |  | REUSABLE | REUSABLE | RECOMMENDED TEMPERATURE:
Fuji |  | SINGLE ENDED | COMBINED WITH CHANNEL BRUSH | EXPOSURE REQUIREMENT:
FUSE |  | DOUBLE ENDED |  | Is the HLD approved for use in the AER?
Are the AER hook-ups the correct part for each scope accessory?
Attach scope inventory list (dropdown)
Yes
Device Used: | Choose one of the following brands: |  |  | No
Scope Buddy | EndoGator
Endo Flush | Pentax
Other | ERBE
Fuji